## Attachment 4
Firm Name:
General Requirements
Unless noted as "desirable," all requirements listed are assumed to be mandatory. Non-compliance with a mandatory requirement will not necessarily eliminate a vendor; however, the number of non-compliant mandatory responses will impact the vendor's overall ranking. All vendors are required to answer the questions to the best of their ability. Some questions are best answered with the drop-down menu options provided: yes, no, yes with configuration, or yes with customization. Other questions will require additional notes.
Number | Requirement | Vendor Response (select an option from the drop down) | Notes
REGULATORY COMPLIANCE
1.1 | ICP system is able to meet all method requirements as specified for the United States Environmental Protection Agency (hereinafter “EPA”) Method 200.7 Revision 4.4. |  |  | changed from "must be" to "is" so they can respond
1.2 | ICP system supplied is able to meet the relevant quality control standards adopted by The NELAC Institute (hereinafter “TNI”) as specified under Rule 64E-1.0015 Florida Administrative Code (hereinafter “FAC”).
HARDWARE
2.1 | Spectrometer is fully computer controlled with background correction capability.
2.2 | Automated computer control of all plasma gas flows, plasma ignition and shut-down, plasma status, radio frequency (hereinafter “RF”) power, and safety interlocks is provided.
2.3 | Optical system and detector(s) is thermostatically controlled for stability.
2.4 | Optical system and detector(s) is purged or sealed with an appropriate inert gas such as argon or nitrogen.
2.5 | Spectrometer utilizes both axial and radial viewing of the plasma.
2.6 | Spectrometer is designed to resolve analytical interferences and limitations associated with axially viewing the plasma.
2.7 | Plasma viewing position is computer controlled with position location electronically storable.
2.8 | Automatic optimization of the viewing position of the plasma is to be provided.
2.9 | RF generator is computer controllable with power settings being electronically storable.
2.1 | RF generator is compliant with all relevant Federal Communications Commission’s regulations.
2.11 | Spectrometer is solid state detector technology utilizing either charge injection or charge coupled device(s).
2.12 | Spectrometer utilizes an internal wavelength reference to ensure optical stability through drift correction.
2.13 | Internal wavelength reference operates automatically at user specified time intervals.
2.14 | Spectrometer has a safety interlocks which continually monitor all required utilities, all gas pressures, plasma condition and plasma compartment access for appropriate operating ranges or status.
2.15 | Safety interlocks is capable of automatic plasma shutdown to avoid damage to either the ICP-AES system or the ICP system operator.
2.16 | Full plasma is safely viewable such that the ICP system operator is not exposed to ultraviolet (UV) or RF emission.
SAMPLE INTRODUCTION SYSTEM
3.1 | Sample introduction system is able to effectively transport samples with acid content of up to 20% and industrial waste samples with complex matrices containing total dissolved solids up to 5%.
3.2 | Plasma torch design allows for easy removal and reinstallation of the plasma torch without the use of tools.
3.3 | Plasma torch design provides simple realignment when reinstalling the plasma torch.
3.4 | Nebulizer gas flow is computer controllable with electronic storage of settings.
3.5 | A peristaltic pump capable of computer controlled variable speed with at least 3 channels is provided.
AUTOSAMPLER
4.1 | Autosampler is fully computer controllable with electronically storable settings.
4.2 | Autosampler is fully integrated into the operation of the ICP system.
4.3 | Autosampler accommodates at least 8 positions for standards (50 mL volume tubes) and at least 100 positions for samples (15 mL volume tubes).
4.4 | Autosampler allows for random access to all autosampler positions.
4.5 | Autosampler components which come in contact with standards and samples is corrosion resistant.
ICP SYSTEM COMPUTER
5.1 | All miscellaneous equipment such as firmware, interface cards and cables necessary to fully control and operate the ICP system is provided.
5.2 | The disk storage capacity is large enough to accommodate data from a minimum of 100 runs of 100 samples and quality control standards.
5.3 | The ICP system computer meets or exceed the following specifications: Intel Core i7 2.8 GHz CPU, 8 GB RAM, 256 GB solid state drive, 22” monitor
SOFTWARE AND OPERATIONAL SPECIFICATIONS
6.1 | ICP system software is compatible with Windows 11 edition operating system.
6.2 | ICP system software is supplied and provides full and integrated control of all system hardware on date of installation.
6.3 | Complete ICP system software packages are supplied for ICP system computer with which the hardware is operable.
6.4 | ICP system software provides multiple calibration techniques including but not limited to single point, linear, linear through zero, and weighted linear regressions.
6.5 | ICP system software is able to store and recall calibration curves.
6.6 | ICP system is able to automatically optimize ICP system operating variables including nebulizer gas flows, RF generator settings, optical position for viewing plasma, sampling time and detector(s) integration time. These variables are individually controllable by the ICP system operator through the ICP system computer.
6.7 | ICP system software has the capability as part of the stored analytical method to increase the peristaltic pump speed to reduce the delivery time for sample transport to the plasma.
6.8 | ICP system is able to measure analyte and background emission spectral data simultaneously.
6.9 | ICP system provides for automatic or manual background correction capability which is ICP system user selectable.
6.1 | All raw spectral data is electronically stored and preserved in its original state.
6.11 | ICP system software provides the capability to reprocess the raw spectral data with ICP system operator changes to peak wavelengths, background correction points, interelement correction factors, and standards.
6.12 | ICP system software is capable of storing all ICP system operating variables and clearly linking this stored information with the data produced using the associated analytical method.
6.13 | ICP system software has the capability to determine and apply interelement correction factors to effectively resolve spectral interferences.
6.14 | ICP system software is able to utilize internal standardization to automatically correct for physical interferences possible with variation in sample nebulization and transport processes.
6.15 | Specification of sample dilutions is storable in ICP system software with automatic correction for these factors.
6.16 | ICP system software allows for the same autosampler location to be used for multiple quality control parameters (ex. ICV with ±5% acceptance limit, CCV with ±10% acceptance limit).
6.17 | ICP system software allows for insertion of samples into an analytical sequence while the sequence is concurrently being analyzed by the ICP system.
6.18 | ICP system software is able to provide spectral displays for all elements being analyzed.
6.19 | ICP system software is able to display calibration curves for all elements being analyzed.
6.2 | ICP system software contains a comprehensive spectral library of emission lines of potential interest and of interfering elements for each wavelength. A number of primary analytical lines should be presented when an element method is accessed.
6.21 | ICP system software allows for data review from previous analytical runs while the ICP system is simultaneously executing sample analysis.
6.22 | ICP system software provides for ICP system operator definable quality control standards and samples which operate using ICP system operator specified acceptance criteria for each analyte wavelength.
6.23 | ICP system software is capable of utilizing ICP system operator specified acceptance criteria for each analyte wavelength to automatically flag unacceptable quality control standards and samples results.
6.24 | Required ICP system software provides response options for quality control failures of ICP system operator specified acceptance criteria for each analyte wavelength and include both “repeat and stop” and “repeat and continue” for all ICP system operator defined quality control standards and samples.
6.25 | ICP system software allows for ICP system operator defined plasma shutdown parameters to include when shutdown will occur such as at the end of an automated analysis or after a quality control failure resulting in a stop condition.
6.26 | ICP system software provides the capability to schedule an ICP system operator defined rinse sequence immediately prior to an automated plasma shutdown.
6.27 | ICP system software is capable of exporting data into a format which is directly importable and compatible with the existing Clinisys Element LIMS.
PERFORMANCE | ICP SYSTEM COMPUTER
7.1 | Detection limits to be provided to evaluate the performance of the ICP system.
7.2 | Detection limits data is determined and calculated based upon the MDL procedure described in 40 CFR Part 136 Appendix B, Revision 2.
7.3 | ICP system has analytical linearity over a minimum of 5 orders of magnitude without switching to alternate emission lines.
7.4 | Firm to provide complete information regarding the specific ICP system operating variables used to produce the detection limit data.
INSTALLATION | Installation
8.1 | At no additional cost to the entity, firm will install the ICP system in the Lake County Water Resource Laboratory (LCWRL) located at 28432 Lady of the Lakes Ave., Tavares, FL 32778.
8.2 | At no additional cost to the entity, firm will provide all relevant information concerning the installation in a documented form at least three (3) weeks prior to the scheduled date of installation.
8.3 | All costs of installation and initial ICP system testing for conformance to specifications at the time of installation shall be included in the bid price.
8.4 | ICP system installation shall be complete, and the instrument made fully operable no later than 60 days after issuance of the purchase order.
TRAINING | Training
9.1 | Firm provides on-site training for LCWA personnel on the operation and user maintenance requirements of the ICP system.
9.2 | This on-site training is a minimum time period of two days and shall occur within three (3) days following the date of the completion of the installation of the ICP system and the initial ICP system testing for conformance to specifications.
9.3 | Firm agrees to allow LCWA personnel to videotape this training for internal purposes.
CONFORMANCE AND WARRANTY
10.1 | Firm agrees to accept the return of the ICP system should the ICP system fail to conform to any required specification within this 30-day period.
10.2 | Should the ICP system be returned to the firm, the firm agrees to provide a full refund of the bid price of the ICP system to the Lake County Water Authority.
10.3 | Firm agrees that, for the first year after acceptance, warranty shall be available .
10.4 | Firm agrees that, for a period of ten (10) years, commencing with the end of the one-year period following installation and acceptance, the firm shall make available one (1) year service maintenance for preventative maintenance and emergency service on both the ICP system hardware and software on a recurring annual basis.
10.5 | During the warranty or maintenance period Contractor shall provide technical support services for a minimum 8:00 to 5:00 PM Eastern Time, Monday through Friday.
DOCUMENTATION
11.1 | The following documents shall be provided separately, in digital format, in proposal submitted.
11.1.1 | Full technical specifications of ICP system
11.1.2 | System manuals for all hardware and software
11.1.3 | Training manuals
11.1.4 | Parts, supplies and accessories (additional features, capabilities or modules) catalog
11.2 | The following documents will be provided after award in paper format:
11.2.1 | System manuals for all hardware and software
11.2.2 | Proof of an established network of service and technical applications support personnel that shall respond promptly and effectively to any problems associated with the operation of the ICP system.
11.2.3 | Provide proposed training schedule
EXISTING EQUIPMENT AND TRADE-IN
12.1 | Firm offers discount(s) from the bid price for the trade-in of an ICP spectrometer system, iCAP 7000 services Thermo. This ICP spectrometer system has been in use since 2020 and has been covered by a manufacturer's service contract since installation. This system is currently in use.
12.2 | Should firm elect to offer a discount from the bid price for the trade-in, bidder must separately list the value of the discount offered on the bid response page.
12.3 | Firm must list a zero (0) trade-in value on the bid response page when bidder offers no discount and does not elect to receive trade-in equipment.
12.4 | The acceptance of any trade-in value offered by the firm is at the sole discretion of the Lake County Water Authority.
12.5 | The trade-in exchange must occur after all conformance and validation testing of the new instrument is completed. Firm is responsible for all packing, labor, and transportation costs involved in acceptance and removal of trade-in equipment.
SERVICE AND WARRANTY
13.1 | For a period of one year following the installation and acceptance, all repairs, parts, labor, travel and living expenses shall be provided by the firm at no additional cost to the LCWRL for as needed to maintain the ICP system proposed.
13.2 | For a period of 9 years commencing with the end of the one-year period following installation and acceptance, the vendor shall make available one-year service maintenance contract for preventative maintenance and emergency service on both the ICP system hardware and software on a reoccurring annual basis with no additional terms and conditions.
13.3 | Software upgrades are included with yearly maintenance for no additional cost.
COSTS
14.1 | Initial and Annual fees shall be disclosed. Any additional costs for services or equipment shall be outlined in Attachment 3, Pricing Sheet.
14.2 | Annual cost, after year two will not exceed 3% above the previous years’ annual license and maintenance cost.

## Please Ignore (hidden)
Internal - Please ignore
Please ignore this worksheet
Vendor Response
Yes
No
Yes with configuration
Yes with customization